chemical sensing
linear devices

---
chemical sensing
introduction to chemical sensing and sensors
  - vapor detection techniques (mostly chemistry)
  - bulk detection techniques (mostly physics)
- in general, “spectroscopic” techniques
- in parallel, algorithmic approaches to “signal-to-symbol transformation” for
  these sorts of signals

---
approaches to chemical sensing
- identify the nuclei, e.g.,
  neutron-activation or -ray spectroscopy
- identify the atoms, e.g.,
  flame emission spectroscopy
  - these are great for, e.g., prospecting for iron ore: you might not care whether you find FeO, Fe2O3, Fe2S3, or any other iron compound, as long as it is Fe
  - but if you want to know, e.g., how a computer works it doesn’t do you a lot of good to grind it up and analyze the dust for H, C, N, O, Si, Al, Fe, Cu, Au, Sn, Pb, etc

---
identify positive and negative ions of atoms, fragments of molecules, most small molecules, some big molecules,
e.g., mass spectrometry:
  - but there are many ways to make the same mass, e.g., H3COCH3 (acetone) and H3CCH2OH (ethyl alcohol) look the same at any practical mass resolution, and both look the same as NO2 and isotopes of Ca, Sc, Ti, and V (all atomic mass 46) at low resolution, i.e., at high detection sensitivity

---
identify effect of molecular solubility (partition) between two solvents on transport time through a “sticky pipe”, e.g., gas and liquid chromatography
  - “retention time” not unique
  - concatenated techniques, e.g.,
    GC-MS, effective but slow and expensive
identify electric-field induced drift rate of molecular ions through a gas, e.g., ion mobility spectrometry (IMS, plasma chromatography, …)
  - airport hand luggage sniffers
    http://www.sensir.com/Smiths/InLabSystems/IonScan/IonScan.htm

---
identify characteristic x-ray spectral attenuation of materials of particular interest in particular places
  - airport “color” x-ray machines for explosives, drugs
and probably a hundred specialized technologies relying on ...
  - photoelectric effect
  - speed of sound
  - infrared absorption
  - etc etc etc ... taking advantage of some unusual chemical or physical property of the specific analyte

---
- in general, we can do quite well these days with complex instruments whose scale is room size or even desk size ... and more recently, desktop monitor size ...
- but there is a demand for low-cost hand-held (or robot-held) equivalents …
- many are based on “chemi-resistors”, “chemi-transistors”, “chemi-capacitors”, etc
  - covered briefly on the white-board recently
first we will discuss “laboratory” chemical analytical instruments and how they are being/might be miniaturized

---
spectroscopies

---
spectroscopies
when a single component produces a mix of separable responses ...
  - example: the optical spectrum of a particular isotope of iron (Fe)
    - electron state transitions between all possible energy levels of the atom (subject to some “selection rules”)
  - example: the ion mass spectrum of a molecule of heptane (gasoline is mostly C7H16)
    - C+, CH+, CH2+, CH3+, CH3C+, CH3CH+, CH3CH2+, CH3CH2C+, CH3CH2CH+, CH3CH2CH2+, CH3CH2CH2C+, ..., CH3CH2CH2CH2CH2CH2CH3+

---
or a mixture produces a complex response for each component
  - can sometimes pre-separate the mixture components
    - gasoline: ..., hexane (C6H14), heptane (C7H16), octane (C8H18), ... can be separated in time domain (e.g., gas chromatography)
(structural separation by MS)
(temporal separation by GC)

---
optical spectroscopy

---
illustrates the general principle …
inevitable tradeoff between your ability to separate spectral components (resolution, selectivity) and your ability to detect small quantities (sensitivity)

---
miniaturization example
Ocean Optics:
optical spectrometer
optics and electronics
on a PC card; separate
light source (below),
and fiber optic (blue)
light input path

---
example: VIS-NIR Diffuse Reflectance Spectrum to Measure Fish Freshness
(probe: light in and out)
(monochromator: specific color light out)

---
mass spectrometry

---
mass spectrometry
- usually a separation based on mass of positive ions; sometimes negative ions, rarely neutrals
- usually all the ions are accelerated to the same energy (and filtered to remove outliers)
- velocity thus depends on mass: v = (2 W/m)1/2
- velocity measured by time-of-flight,
  by trajectory in a magnetic field, etc,
  in many different geometries

---
smaller lower cost alternative:
quadrupole mass spectrometers
  - ions move under combined influence of DC and oscillating (RF) electric fields; most orbits are unbounded, but for any particular mass there is a small region in the DC/RF amplitude plane where they are bounded
    - equations of motion analogous to the inverted pendulum
      - similar to the inverted pendulum application made famous as an example of fuzzy logic control

---
miniaturization example
argon/air/helium, 500 micron diameter rods, 3 cm long
http://www3.imperial.ac.uk/portal/page?_pageid=189,618267&_dad=portallive&_schema=PORTALLIVE

---
chromatographies

---
gas chromatography
- pipe coated (or packed with grains that are coated) with a “sticky” liquid (“stationary phase”)
- inert gas (e.g., He) flows through the pipe (“column”)
- mixture (e.g., gasoline) squirted into “head”
- gas (“mobile phase”) carries it over the liquid
- mixture components move at different effective speeds due to different equilibria between phases
- components emerge at column “tail”
  - detect with a “universal” detector
  - or use as inlet to mass or optical spectrometer, etc

---
miniaturization example
http://eetd.llnl.gov/mtc/Instruments.html
(another instrument – fewer details – link to this one has disappeared)

---
MANY similar techniques:
liquid chromatography
  - liquid mobile phase, solid or liquid stationary phase
ion mobility chromatography
  - ion drift velocity through a gas under influence of an electric field (airport explosives detector principle)
electrophoresis
  - molecules drift through a gel under influence of an electric field (used in many medical tests)
real old fashioned chromatography
  - dye-like chemicals separated by different diffusion speed through a packed powder, e.g., chalk stick,
    or soup dribble on table cloth

---
hybrid techniques

---
hybrid or “tandem” techniques
for routinely detecting and identifying any but the simplest chemical species, hybrid techniques are usually employed …
  - GC – MS
  - pre-concentration – IMS (airport explosives)
  - multiple MS stages with collisional decomposition between stages
  - etc

---
LC MS with high-pressure ionizer etc
note analogy to image processing:
not one magic bullet, but a clever
chain of simple unit operations

---
linearity

---
linearity & superposition
all the techniques discussed today are (nearly) linear in several senses of the word
  - output signal linear in sample concentration
  - response to multiple components present simultaneously is the sum of the responses to the individual components separately
    - i.e., little or no cross-sensitivity
later we will discuss sensors where this is not true, e.g., solids state chemical sensors
  - like the SnO2 chemi-resistors discussed previously
if it is true then simple pattern recognition works

---
unraveling overlapping spectra
(or “signatures”)

---
overlapping spectra of a mixture
absent separation (like GC), given the spectrum of a mixture, how best to unravel its components when the component spectra all overlap?
  - arrange your spectrum library in a rectangular matrix:
    - S1 = {s11, s12, s13, ..., s1n}
      1 = hexane, {1,2,3,...,n} = peak IDs
    - S2 = {s21, s22, s23, ..., s2n}
      2 = octane, {1,2,3,...,n} = same peak IDs
    - ... etc ....
    - Sm = {sm1, sm2, sm3, ..., smn}
      m = Xane, {1,2,3,...,n} = same peak IDs

---
consider the inverse problem: given the concentrations, it is very easy to predict what the combined spectrum will be:
  - C = {c1, c2, c3, ..., cm},
    1 = hexane, 2 = octane, ..., m = Xane
  - S = c1S1 + c2S2 + c3S3 + ... + cmSm
or in matrix notation
s c = S:

---
- if we look at exactly as many spectral peaks as there are components in the mixture then the matrix is square, and it is easy: c = s-1 S
- if we have fewer peaks than components then we are up the creek
  - well, we can establish some constraints ...
- if we have more spectral peaks than components in the mixture then what to do?
- more peaks than components means we have
  “extra data” that we can use to improve the
  precision of our result – a sensor fusion opportunity

---
pseudo-inverse method
the trick is to multiply both sides of the equation by sT:
  - s c = S
    (npeaks * ncomponents) (ncomponents * 1) = (npeaks * 1)
  - sT s c = sTS
    (ncomponents * npeaks) (npeaks * ncomponents) (ncomponents * 1)
    = (ncomponents * npeaks) (npeaks * 1)
  - note that sTs is square, so it (generally) has an inverse

---
 c = (sTs)-1-sTS
(ncomponents * 1) =
(ncomponents * ncomponents)(ncomponents * npeaks) (npeaks * 1)
the calculated component concentrations are optimal: exactly the same as least squares fitting
  - i.e., algebraic least squares fit gives the same result
    as matrix solution using pseudo-inverse formalism
yes, of course, there are degenerate cases where sTs doesn’t actually have an inverse, or calculating it is unstable
  - then you need to use better judgement in deciding which peaks to use!

---
caution ...
c = (sTs)-1sTS is the same as the optimal result you would get if you minimized the sum of the squares of the differences between the components of the data set S and a “predicted” data set S = s c:
  -  = Sum((sc - S)i over all npeaks spectral peaks)
    d /dcj = 0 gives ncomponents simultaneous equations which when you solve them for {c} gives the same result as the pseudo-inverse

---
- but (to keep the notation and discussion simple) I’ve left out something important:
  as in our previous discussion about how to combine multiple measurements that have different associated uncertainties, you need to weight each datum by a reciprocal measure of its uncertainty, e.g., 1/i2
  (in both the least-squares and the pseudo-inverse formulations)
- specific ad hoc weighting schemes are often hard to justify with first-principles arguments

---
exercise
the following table shows the major peaks in the mass spectrum of a mixture of FC-43 and FC-70; you can find their individual spectra at http://www.sisweb.com/index/referenc/mscalibr.htm;
use the “EI Positive Ion ...” data; estimate the fractions of FC-43 and FC-70 in this mixture;
first do a “quick and dirty estimate”, then do it as precisely as you can given the data at your disposal; do you get the best result by using all the data, or might it be better to discard, e.g., data from some of the smaller peaks?

---
note: amu means
“atomic mass units”
(called “daltons”, by chemists and biologists)
all the peaks are
normalized to
the biggest one
(CF3  69 amu)